## Sourcing BOM
Item | Ref Numeral | Part ID | Description | Qty | Material / Spec | Manufacturing Process | Make / Buy | Primary Supplier (India) | Secondary Supplier (Global) | Part Number / SKU | Est. Unit Cost (INR) | Lead Time (Days) | CAD Mass (g) | Functional Notes & Standards
1 | 100 | PRT-001_Cuff_Base | Ergonomic Anatomical Forearm Cuff Chassis | 1 | Biocompatible PETG (ISO 10993 compliant) | FDM / MJF Additive Mfg (100% infill) | Make | Local Chennai 3D Hub / Robu.in | Protolabs / Hubs | CUSTOM-PETG-001 | 650 | 3 | 35.12 | R35 anatomical curvature; M2/M3 brass heat-set inserts; twin 21x4mm strap slots
2 | 102 | PRT-002_Cantilever_Spring | Compliant Variable-Stiffness Cantilever Blade | 1 | C75S / AISI 1095 Spring Steel Strip (48-52 HRC) | Precision Wire EDM / Laser Cutting | Make | Steel City / Local Guindy Shim Stock | Misumi / McMaster-Carr | SHIM-C75S-0.40 | 250 | 4 | 4.25 | 115x12x0.40 mm; polished surface Ra 0.4; deburred edges R0.2; infinite fatigue life SF=2.33
3 | 104 | PRT-003_Root_Clamp_Block | Fixed Encastre Root Clamping Block | 1 | Aluminum Alloy 6061-T6 (Black Anodized) | CNC 3-Axis Milling | Make | Ambattur Industrial Estate CNC Hub | Xometry / Misumi | CNC-AL-003 | 450 | 5 | 4.74 | 18x16x7 mm; 2x M3 counterbore; R1.5 mm relief lead-in radius at bottom exiting edge
4 | 106 | PRT-004_MGN7_Rail | Miniature Linear Guide Rail (80 mm length) | 1 | High-Carbon Bearing Steel (SUJ2, 58-62 HRC) | Precision Ground Linear Motion | Buy | Robu.in (HIWIN Distributor, Pune/Chennai) | HIWIN Global / Misumi | MGN7R-80C | 1150 | 2 | 19.26 | 80x7x4.8 mm; M2 counterbores at 15 mm pitch; 51.5 mm stroke covers 3-8 Hz band
5 | 108 | PRT-005_MGN7C_Block | Miniature Linear Recirculating Ball Carriage | 1 | Stainless Steel 440C / Bearing Steel | Precision Linear Ball Bearing | Buy | Robu.in (HIWIN Distributor) | HIWIN Global / Misumi | MGN7C-1R | 850 | 2 | 14.43 | 22.5x17x6.5 mm; 4x M2 tapped mounting holes (8x12 mm spacing); smooth axial glide
6 | 110 | PRT-006_Fulcrum_Pinch_Bracket | Bifurcated Yoke Fulcrum Pinch Bracket | 1 | Aluminum Alloy 6061-T6 (Clear Anodized) | CNC 3-Axis Milling | Make | Ambattur CNC Job Shop, Chennai | Xometry / Hubs | CNC-AL-006 | 550 | 5 | 6.62 | Mounts to MGN7C block; houses upper and lower Ø2.5mm dowel pins with 0.10mm blade clearance
7 | 112 | PRT-007_Friction_Drag_Unit | Asymmetric Friction Sprag Walking Shoe | 1 | Brass CZ121 + Phosphor Bronze Leaf | CNC Micro-Milling & Spring Stamping | Make | Local Chennai Machine Shop | Misumi / McMaster-Carr | CUSTOM-DRAG-007 | 350 | 4 | 4.13 | 35° inclined sprag geometry; mu_fwd=0.08 vs mu_rev=0.35; dynamic walking rectifier
8 | 114 | PRT-008_Distal_Proof_Mass | Secondary Tuned Inertia Member Block | 1 | Precision Extruded Brass CZ121 (8.44 g/cm3) | CNC 3-Axis Precision Milling | Make | Parrys Corner Brass Suppliers / Ambattur | Misumi / McMaster-Carr | CNC-BRASS-008 | 750 | 4 | 97.72 | 28x22x19.5 mm; 2x M2.5 tapped; hits m2=96.5±1.0g; underside Ø12.2mm damper pocket
9 | 116 | PRT-009_Viscoelastic_Damper | Den Hartog Secondary Shear Damper Puck | 1 | Sorbothane OO-50 (Fallback: Shore A 25 Silicone) | Waterjet / Die Punching | Buy | Polymer Products India / Indiamart | Sorbothane Inc. / McMaster | SORB-OO50-12X5 | 380 | 5 | 0.74 | Ø12.0x5.0 mm; loss factor eta=0.65; sizes secondary damping ratio to optimal zeta=0.168
10 | 118 | PRT-010_Protective_Shroud | Dust & Impact Safety Shroud with Window | 1 | Optical Polycarbonate / Clear PETG | SLA Optical / Thermoforming | Make | Think3D Chennai / Imaginarium | Protolabs / Formlabs | SHROUD-PC-010 | 700 | 3 | 45.65 | 142x40x34 mm; 60x5mm calibration window; accommodates 36mm total vertical clearance
11 | 120 | STRAP-001 | Forearm Dual Retention Straps | 2 | Medical Hook-and-Loop (ISO 10993-5) | Textile / Webbing Cut & Ultrasonic Weld | Buy | Surgiwear India / Amazon.in Medical | Velcro USA / 3M Medical | MED-VEL-20X250 | 120 | 1 | 6 | 20mm width x 250mm length; breathable skin-contact backing; secure dorsal attachment
12 | 124 | PIN-001_Dowel_Pins | Precision Ground Fulcrum Pinch Dowel Pins | 2 | Hardened Tool Steel (ISO 8734 h6, 60 HRC) | Precision Centerless Ground | Buy | Fasteners India / Local Guindy | Misumi / McMaster-Carr | PIN-ISO8734-2.5X14 | 40 | 1 | 1.08 | Ø2.5 h6 x 14.0 mm length; low-friction rolling contact; defined fulcrum node
13 | 128 | SCR-M3-SHCS | Root Clamp Encastre Fasteners | 2 | Stainless Steel A2-70 (ISO 4762) | Cold Forged Socket Head Cap Screw | Buy | TVS Fasteners / Local Industrial Hardware | McMaster-Carr | M3X12-ISO4762 | 15 | 1 | 2.2 | M3 x 12 mm; 1.5d thread engagement into brass inserts in base
14 | 130 | SCR-M2.5-CSK | Distal Proof Mass Retention Screws | 2 | Stainless Steel A2-70 (ISO 10642) | Countersunk Flat Head Screws | Buy | TVS Fasteners / Local Industrial Hardware | McMaster-Carr | M2.5X8-ISO10642 | 15 | 1 | 1.1 | M2.5 x 8 mm countersunk; flush mounting to top face of spring blade
15 | FAST-M2 | SCR-M2-RAIL | MGN7 Rail & Carriage Fasteners | 9 | Stainless Steel A2-70 (ISO 4762) | Hex Socket Head Cap Screws | Buy | TVS Fasteners / Local Industrial Hardware | McMaster-Carr | M2X6-ISO4762 | 30 | 1 | 1.2 | M2 x 6 mm; 5 for rail mounting to base, 4 for bracket mounting to MGN7C block
|  | TOTALS | Total Prototype Cost & Device Mass | 27 | [6 empty] | 0 | Max Lead: 5 Days | 0 | Device Mass: 244.2 g (Budget: 50-300g PASS) ¦ Cost: ₹5,310 (~$64 USD)